## FMEA Form
Design Failure Modes and Effects Analysis (dFMEA)
[6 empty] | Platform: | Painfree Applicator |  |  |  | Participants: |  |  |  | Design Location:
[6 empty] | Product Design: | Activator
[6 empty] | Comment: | Original dFMEA for this product | [7 empty] | dFMEA Revision: | A | Date:
[6 empty] | Facilitated By: | G.Freiberg | Ph:
Date:
Feature ID | Step | Affected | Fit | CPID | Function Input Which component function is being addressed? (Must appear on Function Matrix) | Potential Failure Mode What can go wrong with the process step output? - little y's gone bad - converse of outputs in dFMAP | Potential Failure Effects What is the impact of the failure mode to the function of the device? | SEV | Potential Causes What input went bad and how (root cause of input going bad)? | OCC | Controls What controls exist that prevent or detect either the cause or the FM prior to leaving the function step? | DET | RPN | Actions Recommended What are the actions for reducing the SEV. of the effect, the OCC. of the cause or improving DET.? | Plans / Responsibility What is the target completion date and who is responsible? | DSEV | DOCC | DDET | DRPN
Cap - 12010 | 0.5 | 22004.01 | CL | D | Inner diameter where the Button moves axially within the Cap during activation | The Button resists movement axially within the Cap | Device will not activate | 5 | ID of Cap is smaller than the OD of the Button | 7 | Design Review & Dimensional Analysis | 1 | 35 |  |  | 5 | 3 | 1 | 15
[8 empty] | 5 | OD of Button is larger than the ID of the Cap | 7 |  | 1 | 35
[7 empty] | Device is diffuclt to activate | 3 | The degree of interference between the Cap and the Button | 7 |  | 1 | 21
| 0.6 | 22004-1.8 | CL | D | Inner radius that provides clearance for the button to rest against | Inner radius does not provide clearance | Cannot overcome the safety and Component damage | 5 | Cap Radius is too big there will be an interference between the Cap & Button | 7 | Design Review & Dimensional Analysis | 1 | 35
[8 empty] | 5 | Button Radius is too small there will be an interference between the Cap & Button | 7 | Design Review & Dimensional Analysis | 1 | 35
| 0.7 | 22004-0.8 | CL | D | Guide "times" the Button to the Inner Housing sub-assembly | Button is not "timed" with the IHSA | Cannot fire the device or can fire the device in safe mode | 500 | Cap guide is too short | 5 |  | 1 | 2500
[8 empty] | 500 | Cap guide is too narrow | 5 |  | 1 | 2500
[5 empty] | Provides axial alignment between the Cap and Button | Excessive interference between components | Cannot fire the device | 5 | Cap guide is too tall | 7 |  | 1 | 35
[8 empty] | 5 | Cap guide is too wide | 7 |  | 1 | 35
|  | 32092-0.4 | IF | D | Land that provides a bearing surface for the rotation of the Cap into the “armed” position | Excessive interference between components | Cannot rotate the Cap | 5 | Cap is guide is too tall | 7 |  | 1 | 35
|  | 32092-0.5 | CL | D | Stop that prevents the mechanism from rotating counter-clockwise
|  | 32092-0.6 | CL | D | Lock that aligns and maintains the position of the components in the “armed” position
|  | 32092-0.7 | CL | D | Stop that prevents the mechanism from rotating clockwise past a pre-determined position
|  | 32092-0.8 | IF | D | Lift provides a ramp for the Cap as it lifts onto the bearing surface
| 0.8 | 72093-0.7 | CL | D | Assembly Latch which provides a permanent assembly of the Cap to the Mid Housing and works as the rotation stops for the deactivation of the safety
| 0.9 | 22004-1.2 | CL | D | Inner Diameter sized to provide clearance for the rotation of the safety feature
|  | 72093-0.1 | CL | D
| 1.1 | 72093-0.7 | CL | D | Assembly guide orients the Mid Housing to the Cap during the assembly process
| 1.3 | 22004-1.7 | BE | D | Bearing stop for the Button
Button 22004 | 1 | 32092-1.8 | CL | D | Length of Button stoke safety
| 1.3 | 32092-1.1 | CL | D | Wedge safety that when in the safe position the Fingers can not “jump” over
| 1.6 | 32092-1.2 | SF | D | Wedge profile that creates that functionally separates the fingers to activate the device
Note on D8: dFMEA SF = Slip Fit CL = Clearance IF = Interference Fit BE = Bearing Surface
Note on E8: dFMEA C = issue regarding customer P = issue regarding performance I = issue regarding Interactions D = issue regarding Design
Note on I8: How severe is the effect on the function of the device? 20 = Autoinjector miss fires 10 = Autoinjector does not function 5 = Autoinjection does not perform as expected 1 = No affect
Note on K8: If the failure occurred how often would the failure mode occure: 7 = Will Occur 5 = May Occur 3 = Unlikely to Occur 1 = Will not Occur
Note on M8: How well can you detect the cause or failure mode using the current controls? 7 = Cannot Detect 5 = Difficult to Detect 3 = Maybe able to Detect 1 = Easy to Detect
Note on N8: What is the measure of the process risk related to the Severity, Occurence, and Detection? RPN=SEV x OCC x DET